Transcatheter Valve-in-Valve Implantation Edwards Sapien XT in a Direct Flow Valve after Early Degeneration

Bilge Duran Karaduman Asist. Prof.1, Hüseyin Ayhan Assoc. Prof.1, Serkan Bulguroğlu M.D.2, Telat Keleş Prof.2, Engin Bozkurt Prof.3

1 Atılım University, Faculty of Medicine, Department of Cardiology, Medicana International Ankara Hospital

2 Ankara Yıldırım Beyazıt University, Faculty of Medicine, Department of Cardiology, Ankara City Hospital

3 Medicana International Ankara Hospital, Department of Cardiology

Correspondence Author:

Hüseyin Ayhan

e-mail: huseyinayhan44@yahoo.com

phone: +905058656294

fax: +90 312 287 20 01

Address: Atılım University, Faculty of Medicine, Department of Cardiology, Medicana International Ankara Hospital

Street: Çankaya  Postal code: 06800

City/Country: Ankara/Turkey

ABSTRACT

In recent years, the use of bioprosthetic valve (BPV) has increased significantly with both surgical aortic valve replacement (SAVR) and transcatheter aortic valve implantation (TAVI) due to reasons such as the advantage of not using anticoagulants. Nevertheless major disadvantage of all BPV is the risk of early structural valve deterioration, leading to valve dysfunction, and requires reoperation, which significantly increases the risk of mortality or major morbidity especially after SAVR. There are a limited number of TAV-in-TAV case reports due to TAVI BPV degeneration. In our knowledge this is the second report of TAV-in-TAV implantation wherein a previously implanted transfemoral 25-mm nonmetallic Direct Flow SVD valve treated with ViV TAVI via Edwards Sapien XT.

Key Words: structural valve deterioration, valve-in-valve, TAVI, bioprosthetic valve

INTRODUCTION

Aortic Stenosis (AS) is the most common form of valvular heart disease worldwide and the treatment of choice for severe AS has been surgical aortic valve replacement (SAVR) for most patients with low surgical risks (1,2). In SAVR patients’ valve usually have replaced with either a bioprosthetic or mechanical valve. However, the trend toward greater prevalence of bioprosthetic valve (BPV) use has been dramatic over recent years. In addition, BPV is recommended for patients over 65 years of age for the aortic valve position in the guidelines due to their limited durability. Transcatheter aortic valve implantation (TAVI), using only BPV of all types, has become a less invasive alternative for patients with intermediate or high surgical risk with severe AS, and indications are expanding (1,2). A major disadvantage of all BPV, whether SAVR or TAVI, is the risk of early structural valve deterioration (SVD) or degeneration, leading to valve dysfunction, and requires reoperation, which significantly increases the risk of mortality or major morbidity especially after SAVR.

In patients with prohibitive surgical risk there is evidence for the feasibility of valve-in-valve procedures via a transcatheter approach (3-6). However, there is not enough data in the literature about valve-in-valve transcatheter aortic valve implantation (ViV TAVI) for degenerated TAVI bioprosthetic valve (7). We report a case of TAV-in-TAV implantation wherein a previously implanted transfemoral 25-mm nonmetallic Direct Flow (Direct Flow Medical, Santa Rosa, California) valve treated with ViV TAVI via Edwards Sapien XT (Edwards Lifesciences, Irvine, California).

CASE

Patient was an 81-year-old hypertensive female referred with severe AS and progressive dyspnea New York Heart Association (NYHA) Class III. Three years ago she underwent TAVI with a 25 mm Direct Flow bioprosthesis because of the high surgical risk. Transthoracic echocardiography (TTE) demonstrated aortic valve bioprosthesis stenosis or SVD (effective orifice area (EOA); 0.62 cm2) with a mean gradient of 46 mmHg with mild aortic regurgitation, increased systolic pulmonary artery pressure and a preserved LVEF (55%). More important for ViV procedures is the BPV’s actual internal diameter, which takes into account the internal diameter of the sewing ring, the struts, and the leaflets. Multi-slice computed tomography (MSCT) and transesophageal echocardiography (TEE) were performed to determine the operation feasibility (peripheral arteries, aortic annulus sizing) and procedural technique (to determine BPV stent, ring or calcification). MSCT and TEE showed that degeneration of bioprosthetic valves with extensive calcification without thrombus and vegetation (Figure 1). Although there was no any explanation for this severe calciﬁcation such as renal dysfunction or history of radiation therapy in this short time after the first TAVI. She was re-evaluated by the heart-team and ViV TAVI via a transfemoral approach was recommended (TAV-in-TAV intervention).

The choice of transcatheter heart valve (THV) type, size and the implant depth within the surgical valve are critical to optimize early and late outcomes of ViV. After evaluation we performed a transfemoral TAV-in-TAV implantation under deep sedation using a 23 mm Edwards Sapien XT THV without predilatation under rapid ventricular pacing (Figure 2). But the valve was not expanded enough. Therefore, we decided to do postdilatation with 25 mm aortic balloon (Figure 3). At the final aortography there was no severe paravalvular leak (PVL) (Figure 4) and after the procedure, mean aortic valve gradient decreased from 46 to 14 mm Hg.  There were no complications and patient was discharged 2 days post procedure. At 1st and 6th month follow-up patient now has NYHA class I symptoms with no regurgitation seen on TTE.

DISCUSSION

With the aging population, aortic stenosis is gradually increasing, and BPV is preferred to avoid systemic anticoagulation for valve replacement. Although BPV has less thromboembolic risk, their durability is limited. When structural valve deterioration develops in BPV, redo SAVR has been the gold standard for the treatment of degenerated BPV. However, the risk of redo SAVR is of course not low and a recent study of the Society of Thoracic Surgeons (STS) national database determined a signiﬁcant increase of in-hospital mortality for repeat isolated SAVR compared with ﬁrst-time SAVR (4.6% vs. 2.2%; p < 0.0001) (8). In recent years, ViV-TAVI has become a less invasive option in these patients but the optimal re-intervention to replace a SVD biological prosthesis after SAVR or TAVI remains uncertain. TAVI has been increasingly used in recent years and it is seen that it will be used in low-risk patients with new studies (9-11). This means that TAVI will be done to younger patients and we will see degenerated TAVI valve cases in the coming years. While ViV-TAVI has been featured in the literature for valve BPV degeneration in recent years, the conditions of TAVI BPV degeneration are uncertain. Treatment of transcatheter heart valve degeneration with SAVR is not possible due to these patients already have a high surgical risk. The Edwards SAPIEN XT and Medtronic CoreValve valve have been approved by the FDA for ViV procedures. Although ViV use is frequent in the literature due to acute valve embolization or device failure, TAV-in-TAV intervention is very rare in the literature due to TAVI valve degeneration. Butter et al., similar to our case, they treated degeneration of the direct flow valve with TAVI (7). In this case report; 76 year old woman was treated with a transfemoral 25-mm nonmetallic Direct Flow valve 17 months ago. After than the patient was admitted with recurrent aortic stenosis and worsening symptoms and they detected calciﬁed conglomerate without any leaﬂet motion in the TEE. In cases of BPV deterioration, infective endocarditis, vegetation, thrombus and dehiscence need to be excluded, as in our case. In this case, Edwards Sapien XT was used like us, but we used 23 mm and they used 20 mm. Even in our case, postdilation was required since there was not enough opening after ViV. The Direct flow system, with its metallic backbone, inflatable and deflatable structure, allows precise placement to the annulus, repositioning, evaluation of the valve performance in its final position and retrieval if necessary. Acceptable data are available regarding the durability of some TAVI valves, and at least surgical bioprosthetic valves are durable. However the studies on Direct flow system, which is one of the 2nd generation TAVI systems, do not have long-term results (12-15). In these studies, it is generally seen that DF has similar acute and short-term results with other first and second generation valves. TAV-in-TAV implantation is possible due to device failure, acute aortic regurgitation or embolization during the procedure, and there are satisfactory short- and mid-term results in the literature (16-17-18). However, TAV-in-TAV implantation is very rare in the literature due to TAVI bioprosthesis valve structural deterioration. Direct flow valve degeneration occurred within 17 months in the study of Butter et al. and within 3 years in our study. In our case, it is seen that the valve is incorrectly placed and one side is in LVOT instead of annulus, which in turn disrupts the current dynamics. In our case, it is seen that the valve was implanted incorrectly and one side of valve was in left ventricle outflow tract instead of aortic annulus, thereby disrupting the flow dynamics. Also, due to the solutions used in the preparation phase of the DF valve, the feature of preventing calcification may not be added.

CONCLUSION

To the best of our knowledge, this is the second report of TAV-in-TAV implantation of a balloon expandable Edwards Sapien XT valve in a Direct Flow valve.

Disclosures

Engin Bozkurt is a proctor for Edwards Lifesciences. The other investigators have no conflicts of interest to declare.

Figure Legends

Figure 1: Multi-slice computed tomography and Transesophageal echocardiography showed that structural valve deterioration of Direct Flow bioprosthetic valves

Figure 2: Angiographic images of TAV-in-TAV implantation with 23 mm Edwards Sapien XT

Figure 3: Postdilatation with 25 mm aortic balloon after TAV-in-TAV implantation

Figure 4: Final aortography

REFERENCES

- Baumgartner h, Falk V, Bax JJ, et al. 2017 ESC/EACTS Guidelines for the management of valvular heart disease. Eur Heart J 2017;38:2739–91.

- Nishimura, R. A., Otto CM, Bonow RO, et al. 2017 AHA/ACC focused update of the 2014 AHA/ACC guideline for the management of patients with valvular heart disease: a report of the American College of Cardiology/American Heart Association task force on clinical practice guidelines. J. Am. Coll. Cardiol. 70, 252–289 (2017).

- Ferrari E, Stortecky S, Heg D, et al. The hospital results and 1-year outcomes of transcatheter aortic valve-in-valve procedures and transcatheter aortic valve implantations in the native valves: the results from the Swiss-TAVI Registry. Eur J Cardiothorac Surg 2019;56:55–63.

- Dvir D, Webb JG, Bleiziffer S, et al. Transcatheter aortic valve implantation in failed bioprosthetic surgical valves. JAMA 2014;312:162–70.

- Webb JG, Murdoch DJ, Alu MC, et al. 3-Year outcomes after valve-in-valve transcatheter aortic valve replacement for degenerated bioprostheses: the PARTNER 2 registry. J Am Coll Cardiol 2019;73:2647–55.

- Tam, Derrick Y., Vo TX, Wijeysundera HC, et al. Transcatheter valve‐in‐valve versus redo surgical aortic valve replacement for the treatment of degenerated bioprosthetic aortic valve: A systematic review and meta‐analysis. Catheterization and Cardiovascular Interventions, 2018, 92.7: 1404-1411.

- Butter C1, Tambor G2, Neuss M2, et al. First successful transfemoral implantation of an Edwards Sapien XT valve in a direct flow valve after early restenosis. JACC: Cardiovascular Interventions, 2016, 9.1: e1-e2.

- Kaneko T, Vassileva CM, Englum B, et al. Contemporary outcomes of repeat aortic valve replacement: a benchmark for transcatheter valve-in-valve procedures. Ann Thorac Surg 2015;100:1298–304.

- Butter C1, Tambor G2, Neuss M2, et al. Transcatheter versus surgical aortic valve replacement in patients with severe aortic valve stenosis: 1-year results from the all- comers NOTION randomized clinical trial. J. Am. Coll. Cardiol. 65, 2184–2194 (2015).

- Popma JJ, Deeb GM, Yakubov SJ, et al. Transcatheter aortic-valve replacement with a self-expanding valve in low-risk patients. New England Journal of Medicine, 2019, 380.18: 1706-1715.

- Mack MJ1, Leon MB1, Thourani VH1, et al. Transcatheter aortic-v alve replacement with a balloon- expandable valve in low- risk patients.  N. Engl. J. Med. 380, 1695–1705 (2019).

- Schofer J, Colombo A, Klugmann S, et al.. Prospective multicenter evaluation of the direct flow medical transcatheter aortic valve. J Am Coll Cardiol. 2014;63:763-8.

- De Marco F, Latib A. Tools and Techniques - Clinical: the inner curve technique for implantation of the Direct Flow Medical® transcatheter aortic valve. EuroIntervention. 2014;10:400-2.

- Naber CK, Pyxaras SA, Ince H, et al. Real-world multicentre experience with the Direct Flow Medical repositionable and retrievable transcatheter aortic valve implantation system for the treatment of high-risk patients with severe aortic stenosis. EuroIntervention. 2016;11:e1314-20.

- Naber CK1, Pyxaras SA, Ince H, et al. A multicentre European registry to evaluate the Direct Flow Medical transcatheter aortic valve system for the treatment of patients with severe aortic stenosis. EuroIntervention: journal of EuroPCR in collaboration with the Working Group on Interventional Cardiology of the European Society of Cardiology, 2016, 12.11: e1413-e1419.

- S. Toggweiler, D. A. Wood, J. Rodés-Cabau et al., “Transcatheter valve-in-valve implantation for failed balloon-expandable transcatheter aortic valves,” JACC: Cardiovascular Interventions, vol. 5, no. 5, pp. 571–577, 2012.

- R. R. Makkar, H. Jilaihawi, T. Chakravarty et al., “Determinants and outcomes of acute transcatheter valve-in-valve therapy or embolization: a study of multiple valve implants in the U.S. PARTNER trial (Placement of AoRTic TraNscathetER Valve Trial Edwards SAPIEN Transcatheter Heart Valve),” Journal of the American College of Cardiology, vol. 62, no. 5, pp. 418–430, 2013.

- Tzimas G1, Eeckhout E1, Antiochos P1, et al. Percutaneous Valvular Closure Followed by TAV-in-TAV Intervention during a Single Procedure in order to Treat a Severe Paravalvular Leak after Performing TAVI in a Bicuspid Aortic Stenosis. Case reports in cardiology, 2019.